## STEP BY STEP GUIDE
PIC32CXMTx Demo Board Calibration Spreadsheet
STEP 1: CONFIGURE THE INPUT SOURCES
Type the values in the yellow cells:
PARAMETER | VALUE
Test Current | 30
Test Voltage | 220
Meter Tester Freq | 50
Meter Tester I-V Phase | 60
Meter Tester V-V Phase | 120
STEP 2: CONFIGURE THE CALIBRATION COMPUTATION SETTINGS AND THE CHANNELS TO BE CALIBRATED
Write "1" or "0" in the yellow cell to enable or disable to voltage to voltage phase correction
SETTING | VALUE (0=NO; 1=YES)
V-V PHASE CORRECTION | 1
Write "1" or "0" in the yellow cell to enable or disable the channel to be calibrated. The blue cells contain suggested values for the common configuration of the devices supported.
CHANNEL | VALUE (0=NO; 1=YES) | PIC32CXMTSH | PIC32CXMTC + ATSENSE301 | PIC32CXMTC + MCP3912 | PIC32CXMTC + MCP3913 | PIC32CXMTC + MCP3914
IA | 1 | 1 | 1 | 1 | 1 | 1
VA | 1 | 1 | 1 | 1 | 1 | 1
IB | 1 | 1 | 1 | 1 | 1 | 1
VB | 1 | 1 | 1 | 1 | 1 | 1
IC | 1 | 0 | 1 | 0 | 1 | 1
VC | 1 | 0 | 1 | 0 | 1 | 1
IN | 0 | 0 | 0 | 0 | 0 | 0
STEP 3 (OPTIONAL): SET THE CALIBRATION CONSTANTS TO THE DEFAULT VALUES (UNITY GAIN AND ZERO PHASE ADJUSTMENTS)
This is an optional step, because the calibration parameters can be calculated using parameterts not having the default values
Using the metrology console, send the default calibration constants to the board. Please check the board documentation showing how to work with the terminal console
Send the commands to the board, copying from the green/purple cells and pasting to the terminal console
Use these commands to send the calibration registers one by one:
REGISTER | UNITY GAIN COMMAND | REGISTER | ZERO PHASE COMMAND
CAL_M_IA | DCW[52](20000000) | CAL_PH_IA | DCW[60](0)
CAL_M_VA | DCW[53](20000000) | CAL_PH_VA | DCW[61](0)
CAL_M_IB | DCW[54](20000000) | CAL_PH_IB | DCW[62](0)
CAL_M_VB | DCW[55](20000000) | CAL_PH_VB | DCW[63](0)
CAL_M_IC | DCW[56](20000000) | CAL_PH_IC | DCW[64](0)
CAL_M_VC | DCW[57](20000000) | CAL_PH_VC | DCW[65](0)
CAL_M_IN | DCW[58](20000000) | CAL_PH_IN | DCW[66](0)
Or use these commands to send the calibration registers as groups:
REGISTER SET | UNITY GAIN COMMAND / ZERO PHASE COMMAND
CAL_M_x | DCM( 52:20000000; 53:20000000; 54:20000000; 55:20000000; 56:20000000; 57:20000000; 58:20000000 )
CAL_PH_x | DCM( 60:0; 61:0; 62:0; 63:0; 64:0; 65:0; 66:0 )
STEP 4: READ THE MEASUREMENTS FROM THE BOARD
Send to the board the commands in dark blue cells, and Copy and Paste ONLY THOSE EXACT LINES from the serial port interface results (soft blue).
>par u
Present voltage is :
Ua=217.920V Ub=217.897V Uc=217.947V
>par i
Present current is :
Ia=29.8207A Ib=29.7977A Ic=29.7782A Ini=0.0000A Inm=0.0000A Inmi=0.0000A
>par a
Voltage and current angle is :
Angle_A=59.959 Angle_B=59.951 Angle_C=59.945 Angle_N=0.000 Angle_VAB=119.997 Angle_VAC=-119.998
>dcr
00 STATE_CTRL 01 FEATURE_CTRL 02 AFE_SELECTION 03 CHANNEL_MATRIX
2 40000700 4 FF420531
04 HARMONIC_CTRL 05 METER_TYPE 06 M 07 N_MAX
0 CCCC 0 1130
08 PULSE0_CTRL 09 PULSE0_K_t 10 PULSE1_CTRL 11 PULSE1_K_t
810001D0 500000 810201D0 500000
12 PULSE2_CTRL 13 PULSE2_K_t 14 SYNTH_ADDR 15 CREEP_THR_P
1D0 500000 0 2E9A
16 CREEP_THR_P_A 17 CREEP_THR_P_B 18 CREEP_THR_P_C 19 CREEP_THR_Q
2E9A 2E9A 2E9A 2E9A
20 CREEP_THR_Q_A 21 CREEP_THR_Q_B 22 CREEP_THR_Q_C 23 CREEP_THR_I
2E9A 2E9A 2E9A 212D
24 CREEP_THR_I_A 25 CREEP_THR_I_B 26 CREEP_THR_I_C 27 CREEP_THR_S
212D 212D 212D 2E9A
28 PWR_OFFS_CTRL 29 PWR_OFFS_P 30 PWR_OFFS_Q 31 PWR_OFFS_S
88000000 FFFFFF8A 378 0
32 SWELL_THR_VA 33 SWELL_THR_VB 34 SWELL_THR_VC 35 SAG_THR_VA
5E84F62 5E84F62 5E84F62 1A2EC26
36 SAG_THR_VB 37 SAG_THR_VC 38 INT_THR_VA 39 INT_THR_VB
1A2EC26 1A2EC26 68BB0A 68BB0A
40 INT_THR_VC 41 RESERVED_C41 42 RESERVED_C42 43 RESERVED_C43
68BB0A 0 0 0
44 K_IA 45 K_VA 46 K_IB 47 K_VB
9A523 19CC00 9A523 19CC00
48 K_IC 49 K_VC 50 K_IN 51 K_VD
9A523 19CC00 0 0
52 CAL_M_IA 53 CAL_M_VA 54 CAL_M_IB 55 CAL_M_VB
20000000 20000000 20000000 20000000
56 CAL_M_IC 57 CAL_M_VC 58 CAL_M_IN 59 CAL_M_VD
20000000 20000000 20000000 20000000
60 CAL_PH_IA 61 CAL_PH_VA 62 CAL_PH_IB 63 CAL_PH_VB
0 0 0 0
64 CAL_PH_IC 65 CAL_PH_VC 66 CAL_PH_IN 67 RESERVED_C67
0 0 0 0
68 CPTR_CTRL 69 CPTR_BUFF_SIZE 70 CPTR_ADDR 71 RESERVED_C71
0 FA0 20000000 0
72 RESERVED_C72 73 RESERVED_C73 74 AFE_CTRL 75 RESERVED_C75
0 0 E000000 0
76 RESERVED_C76 77 RESERVED_C77 78 PWR_OFFS_P_A 79 PWR_OFFS_P_B
0 0 0 0
80 PWR_OFFS_P_C 81 PWR_OFFS_Q_A 82 PWR_OFFS_Q_B 83 PWR_OFFS_Q_C
0 0 0 0
>dsr
00 VERSION 01 STATUS 02 STATE_FLAG 03 CAPTURE_STATUS
30400 8 27 0
04 INTERVAL_NUM 05 N 06 PH_OFFSET 07 FREQ
1D4000F FA0 FF0A3D71 31FFB
08 FREQ_VA 09 FREQ_VB 10 FREQ_VC 11 RESERVED_S11
31FFB 31FFB 31FFB 0
12 TEMPERATURE 13 V_A_MAX 14 V_B_MAX 15 V_C_MAX
0 6093534 608EDAA 6090CD0
16 RESERVED_S16 17 I_A_MAX 18 I_B_MAX 19 I_C_MAX
0 233A272 23434FF 2371896
20 I_Nm_MAX 21 I_Ni_MAX 22 FEATURES 23 N_CYCLE
0 58956 807F6E00 50007
24 RESERVED_S24 25 RESERVED_S25 26 PULSE0_COUNTER 27 PULSE1_COUNTER
0 0 F39 1A65
28 PULSE2_COUNTER 29 RESERVED_S29 30 ZC_N_VA 31 ZC_N_VB
0 0 54F589 51A02B
32 ZC_N_VC 33 RESERVED_S33 34 ATCAL_41_44 35 ATCAL_45_48
584AE9 0 0 0
>dar
00 V_A 01 V_B 02 V_C 03 V_D
45B051F9B426 45AC8C579F3D 45B4A1DD4115 0
04 V_A_F 05 V_B_F 06 V_C_F 07 V_D_F
45B0507D2D77 45ABB28F59C6 45B4AC4E3786 0
08 V_AB 09 V_BC 10 V_CA 11 V_AB_F
D10A26086387 D117E9ADE66F D1125A16C4E0 D1088CA58502
12 V_BC_F 13 V_CA_F 14 RESERVED_A14 15 RESERVED_A15
D11647A57B48 D11339546E7E 0 0
16 RESERVED_A16 17 I_A 18 I_B 19 I_C
0 955CCF3945A 9522B8C5887 94F209D5691
20 I_Ni 21 I_Nm 22 I_A_F 23 I_B_F
19873E3 0 955C97B4085 9522BE56234
24 I_C_F 25 I_Nmi 26 I_Nm_F 27 RESERVED_A27
94F08974F04 19873CD 0 0
28 RESERVED_A28 29 RESERVED_A29 30 RESERVED_A30 31 RESERVED_A31
0 0 0 0
32 P_A 33 P_B 34 P_C 35 P_A_F
CC50E80B33E CC2402D57B4 CC1CE955663 CC52E213F6B
36 P_B_F 37 P_C_F 38 P_N 39 P_N_F
CC210FC5D1C CC1A7B8527A 0 0
40 RESERVED_A40 41 Q_A 42 Q_B 43 Q_C
0 16143F5C38B6 160F38CA4BFF 160CA21C0654
44 Q_A_F 45 Q_B_F 46 Q_C_F 47 Q_N
1614065D2D67 160F5AF6DBB8 160C8875505D 0
48 Q_N_F 49 RESERVED_A49 50 ACC_T0 51 ACC_T1
0 0 130F49B1E3D 210BB84F35F
52 ACC_T2 53 RESERVED_A53 54 RESERVED_A54
0 0 0
>dpcar |  |  |  | (*) Per-cycle measurements, needed for the calibration computations, but useful for informative purposes
00 CYCLE_V_A 01 CYCLE_V_B 02 CYCLE_V_C 03 CYCLE_V_A_F
164BFA1E2CD 164D01DDDB2 164E36BBB61 164BB28D2B0
04 CYCLE_V_B_F 05 CYCLE_V_C_F 06 CYCLE_V_AB 07 CYCLE_V_BC
164CE022081 164DF394B25 42E488AE290 42E95DB261C
08 CYCLE_V_CA 09 CYCLE_V_AB_F 10 CYCLE_V_BC_F 11 CYCLE_V_CA_F
42E7AA4B434 42E3B33443E 42E8DE15527 42E7017EA40
12 CYCLE_I_A 13 CYCLE_I_B 14 CYCLE_I_C 15 CYCLE_I_Nm
2FCB8D512F 2FB9A9996D 2FAAB8E613 0
16 CYCLE_I_A_F 17 CYCLE_I_B_F 18 CYCLE_I_C_F 19 CYCLE_I_Nm_F
2FCB30DC22 2FB997C042 2FA9970352 0
20 CYCLE_P_A 21 CYCLE_P_B 22 CYCLE_P_C 23 CYCLE_P_A_F
41640442B4 41547F30D5 415274C285 4164C4F0AF
24 CYCLE_P_B_F 25 CYCLE_P_C_F 26 CYCLE_P_N 27 CYCLE_P_N_F
4153BF0180 4151FD583D 0 0
28 CYCLE_Q_A 29 CYCLE_Q_B 30 CYCLE_Q_C 31 CYCLE_Q_A_F
7106F9F3EE 70F6CA88D6 70E41237D6 710559FC5B
32 CYCLE_Q_B_F 33 CYCLE_Q_C_F 34 CYCLE_Q_N 35 CYCLE_Q_N_F
70F73C3A95 70E236E079 0
STEP 5: SEND THE COMPUTED CALIBRATION REGISTERS TO THE BOARD
Using the metrology console, send the computed calibration constants to the board.
Send the commands to the board, copying from the green/purple cells and pasting to the terminal console
Use these commands to send the calibration registers one by one:
REGISTER | GAIN COMMAND | REGISTER | PHASE COMMAND
CAL_M_IA | DCW[52](20314A3F) | CAL_PH_IA | DCW[60](FFF611A8)
CAL_M_VA | DCW[53](204E2A66) | CAL_PH_VA | DCW[61](0)
CAL_M_IB | DCW[54](20378E9C) | CAL_PH_IB | DCW[62](FFF628F3)
CAL_M_VB | DCW[55](204F0A40) | CAL_PH_VB | DCW[63](FFFF67BA)
CAL_M_IC | DCW[56](203CD20D) | CAL_PH_IC | DCW[64](FFF6733C)
CAL_M_VC | DCW[57](204D2A95) | CAL_PH_VC | DCW[65](1B95F)
CAL_M_IN | () | CAL_PH_IN | ()
Or use these commands to send the calibration registers as groups:
REGISTER SET | GAIN COMMAND / PHASE COMMAND
CAL_M_x | DCM(52:20314A3F; 53:204E2A66; 54:20378E9C; 55:204F0A40; 56:203CD20D; 57:204D2A95) | [5 empty] | DCM(52:20314A3F; 53:204E2A66; 54:20378E9C; 55:204F0A40; 56:203CD20D; 57:204D2A95; )
CAL_PH_x | DCM(60:FFF611A8; 61:0; 62:FFF628F3; 63:FFFF67BA; 64:FFF6733C; 65:1B95F) | [5 empty] | DCM(60:FFF611A8; 61:0; 62:FFF628F3; 63:FFFF67BA; 64:FFF6733C; 65:1B95F; )
STEP 6: CHECK THE ACCURACY OF THE MEASUREMENTS
Check the accuracy of the measurements. If the result are not as good as expected, check the input sources, check the configuration and come back to step 4

## CALCULATIONS
DO NOT WRITE IN THIS COLUMN IT CONTAINS A COPY OF THE VALUES ENTERED IN "STEP BY STEP GUIDE" SHEET |  |  | Test Values
|  |  | 30 | Test Current |  | Use these CAL constants to set to unity gain with zero phase adjustments.
|  |  | 220 | Test Voltage
|  |  | 50 | Meter Tester Freq |  |  | Unity GAIN | Zero Phase
>par u |  |  | 60 | Meter Tester I-V Phase |  | CAL_M_IA | dcw[52](20000000) | dcw[60](0)
Present voltage is : |  |  | 120 | Meter Tester V-V Phase |  | CAL_M_VA | dcw[53](20000000) | dcw[61](0)
Ua=217.920V Ub=217.897V Uc=217.947V | [5 empty] | CAL_M_IB | dcw[54](20000000) | dcw[62](0)
>par i | [5 empty] | CAL_M_VB | dcw[55](20000000) | dcw[63](0)
Present current is : | [5 empty] | CAL_M_IC | dcw[56](20000000) | dcw[64](0)
Ia=29.8207A Ib=29.7977A Ic=29.7782A Ini=0.0000A Inm=0.0000A Inmi=0.0000A | [5 empty] | CAL_M_VC | dcw[57](20000000) | dcw[65](0)
>par a | [5 empty] | CAL_M_IN | dcw[58](20000000) | dcw[66](0)
Voltage and current angle is :
Angle_A=59.959 Angle_B=59.951 Angle_C=59.945 Angle_N=0.000 Angle_VAB=119.997 Angle_VAC=-119.998 |  |  | DEFAULT
>dcr |  |  | DCM( 52:20000000; 53:20000000; 54:20000000; 55:20000000; 56:20000000; 57:20000000; 58:20000000 ) | [9 empty] | CONTROL REGISTERS
00 STATE_CTRL 01 FEATURE_CTRL 02 AFE_SELECTION 03 CHANNEL_MATRIX |  |  | DCM( 60:0; 61:0; 62:0; 63:0; 64:0; 65:0 ) | [9 empty] | 16 | 1 | 00 | STATE_CTRL | 2 | uQ32.0 | 2
2 40000700 4 FF420531 |  |  |  |  | Format separator |  |  | : |  |  |  |  | 16 | 2 | 01 | FEATURE_CTRL | 40000700 | uQ32.0 | 1073743616
04 HARMONIC_CTRL 05 METER_TYPE 06 M 07 N_MAX |  |  |  | 1 | Use V-V phase correction( 0=NO; 1=YES ) | [7 empty] | 16 | 3 | 02 | AFE_SELECTION | 4 | uQ32.0 | 4
0 CCCC 0 1130 | before (measured) | after (predicted) |  | old cal (hex) | old_CAL(60Hz) | f_test delta_CAL(50Hz) | f_normalized delta_CAL(60Hz) | New_CAL |  |  | raw value | cal commands | 16 | 4 | 03 | CHANNEL_MATRIX | FF420531 | uQ32.0 | 4282516785
08 PULSE0_CTRL 09 PULSE0_K_t 10 PULSE1_CTRL 11 PULSE1_K_t | 29.8206 | 30 | CAL_M_IA | 20000000 | 1 |  | 1.006 | 1.006 | CAL_M_IA | 52 | 20314A3F | DCW[52](20314A3F) | 18 | 1 | 04 | HARMONIC_CTRL | 0 | uQ32.0 | 0
810001D0 500000 810201D0 500000 | 217.9207 | 220 | CAL_M_VA | 20000000 | 1 |  | 1.0095 | 1.0095 | CAL_M_VA | 53 | 204E2A66 | DCW[53](204E2A66) | 18 | 2 | 05 | METER_TYPE | CCCC | uQ32.0 | 52428
12 PULSE2_CTRL 13 PULSE2_K_t 14 SYNTH_ADDR 15 CREEP_THR_P | 29.7979 | 30 | CAL_M_IB | 20000000 | 1 |  | 1.0068 | 1.0068 | CAL_M_IB | 54 | 20378E9C | DCW[54](20378E9C) | 18 | 3 | 06 | M | 0 | uQ32.0 | 0
1D0 500000 0 2E9A | 217.8976 | 220 | CAL_M_VB | 20000000 | 1 |  | 1.0096 | 1.0096 | CAL_M_VB | 55 | 204F0A40 | DCW[55](204F0A40) | 18 | 4 | 07 | N_MAX | 1130 | uQ32.0 | 4400
16 CREEP_THR_P_A 17 CREEP_THR_P_B 18 CREEP_THR_P_C 19 CREEP_THR_Q | 29.7789 | 30 | CAL_M_IC | 20000000 | 1 |  | 1.0074 | 1.0074 | CAL_M_IC | 56 | 203CD20D | DCW[56](203CD20D) | 20 | 1 | 08 | PULSE0_CTRL | 810001D0 | uQ32.0 | 2164261328
2E9A 2E9A 2E9A 2E9A | 217.947 | 220 | CAL_M_VC | 20000000 | 1 |  | 1.0094 | 1.0094 | CAL_M_VC | 57 | 204D2A95 | DCW[57](204D2A95) | 20 | 2 | 09 | PULSE0_K_t | 500000 | uQ8.24 | 0.3125
20 CREEP_THR_Q_A 21 CREEP_THR_Q_B 22 CREEP_THR_Q_C 23 CREEP_THR_I | 0 | #DIV/0! | CAL_M_IN | 20000000 | 1 |  | #DIV/0! | #DIV/0! | CAL_M_IN | 58 | #DIV/0! | #DIV/0! | 20 | 3 | 10 | PULSE1_CTRL | 810201D0 | uQ32.0 | -2130574896
2E9A 2E9A 2E9A 212D |  |  | CAL_PH_IA | 0 | 0 | -0.045461 | -0.054553 | -0.054553 | CAL_PH_IA | 60 | FFF611A8 | DCW[60](FFF611A8) | 20 | 4 | 11 | PULSE1_K_t | 500000 | uQ8.24 | 0.3125
24 CREEP_THR_I_A 25 CREEP_THR_I_B 26 CREEP_THR_I_C 27 CREEP_THR_S |  |  | CAL_PH_VA | 0 | 0 |  | 0 | 0 | CAL_PH_VA | 61 | 0 | DCW[61](0) | 22 | 1 | 12 | PULSE2_CTRL | 1D0 | uQ32.0 | 464
212D 212D 212D 2E9A |  |  | CAL_PH_IB | 0 | 0 | -0.042321 | -0.050786 | -0.054053 | CAL_PH_IB | 62 | FFF628F3 | DCW[62](FFF628F3) | 22 | 2 | 13 | PULSE2_K_t | 500000 | uQ8.24 | 0.3125
28 PWR_OFFS_CTRL 29 PWR_OFFS_P 30 PWR_OFFS_Q 31 PWR_OFFS_S |  |  | CAL_PH_VB | 0 | 0 |  | -0.0032675 | -0.0032675 | CAL_PH_VB | 63 | FFFF67BA | DCW[63](FFFF67BA) | 22 | 3 | 14 | SYNTH_ADDR | 0 | uQ32.0 | 0
88000000 FFFFFF8A 378 0 |  |  | CAL_PH_IC | 0 | 0 | -0.051608 | -0.06193 | -0.052459 | CAL_PH_IC | 64 | FFF6733C | DCW[64](FFF6733C) | 22 | 4 | 15 | CREEP_THR_P | 2E9A | uQ2.30 | 1.1111e-05
32 SWELL_THR_VA 33 SWELL_THR_VB 34 SWELL_THR_VC 35 SAG_THR_VA |  |  | CAL_PH_VC | 0 | 0 |  | 0.0094708 | 0.0094708 | CAL_PH_VC | 65 | 1B95F | DCW[65](1B95F) | 24 | 1 | 16 | CREEP_THR_P_A | 2E9A | uQ12.20 | 0.011377
5E84F62 5E84F62 5E84F62 1A2EC26 |  |  | CAL_PH_IN | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | CAL_PH_VC | 66 | #DIV/0! | #DIV/0! | 24 | 2 | 17 | CREEP_THR_P_B | 2E9A | uQ12.20 | 0.011377
36 SAG_THR_VB 37 SAG_THR_VC 38 INT_THR_VA 39 INT_THR_VB | [12 empty] | 24 | 3 | 18 | CREEP_THR_P_C | 2E9A | uQ12.20 | 0.011377
1A2EC26 1A2EC26 68BB0A 68BB0A | [12 empty] | 24 | 4 | 19 | CREEP_THR_Q | 2E9A | uQ2.30 | 1.1111e-05
40 INT_THR_VC 41 RESERVED_C41 42 RESERVED_C42 43 RESERVED_C43 | [12 empty] | 26 | 1 | 20 | CREEP_THR_Q_A | 2E9A | uQ12.20 | 0.011377
68BB0A 0 0 0 | [12 empty] | 26 | 2 | 21 | CREEP_THR_Q_B | 2E9A | uQ12.20 | 0.011377
44 K_IA 45 K_VA 46 K_IB 47 K_VB | [12 empty] | 26 | 3 | 22 | CREEP_THR_Q_C | 2E9A | uQ12.20 | 0.011377
9A523 19CC00 9A523 19CC00 | [12 empty] | 26 | 4 | 23 | CREEP_THR_I | 212D | uQ12.20 | 0.0080996
48 K_IC 49 K_VC 50 K_IN 51 K_VD | [12 empty] | 28 | 1 | 24 | CREEP_THR_I_A | 212D | uQ12.20 | 0.0080996
9A523 19CC00 0 0 | [12 empty] | 28 | 2 | 25 | CREEP_THR_I_B | 212D | uQ12.20 | 0.0080996
52 CAL_M_IA 53 CAL_M_VA 54 CAL_M_IB 55 CAL_M_VB | [12 empty] | 28 | 3 | 26 | CREEP_THR_I_C | 212D | uQ12.20 | 0.0080996
20000000 20000000 20000000 20000000 | [12 empty] | 28 | 4 | 27 | CREEP_THR_S | 2E9A | uQ2.30 | 1.1111e-05
56 CAL_M_IC 57 CAL_M_VC 58 CAL_M_IN 59 CAL_M_VD | [12 empty] | 30 | 1 | 28 | PWR_OFFS_CTRL | 88000000 | uQ32.0 | -2013265920
20000000 20000000 20000000 20000000 | [12 empty] | 30 | 2 | 29 | PWR_OFFS_P | FFFFFF8A | sQ1.30 | -1.099e-07
60 CAL_PH_IA 61 CAL_PH_VA 62 CAL_PH_IB 63 CAL_PH_VB | [12 empty] | 30 | 3 | 30 | PWR_OFFS_Q | 378 | sQ1.30 | 8.2701e-07
0 0 0 0 | [12 empty] | 30 | 4 | 31 | PWR_OFFS_S | 0 | sQ1.30 | 0
64 CAL_PH_IC 65 CAL_PH_VC 66 CAL_PH_IN 67 RESERVED_C67 | [12 empty] | 32 | 1 | 32 | SWELL_THR_VA | 5E84F62 | uQ0.32 | 0.023076
0 0 0 0 | [12 empty] | 32 | 2 | 33 | SWELL_THR_VB | 5E84F62 | uQ0.32 | 0.023076
68 CPTR_CTRL 69 CPTR_BUFF_SIZE 70 CPTR_ADDR 71 RESERVED_C71 | [12 empty] | 32 | 3 | 34 | SWELL_THR_VC | 5E84F62 | uQ0.32 | 0.023076
0 FA0 20000000 0 | [12 empty] | 32 | 4 | 35 | SAG_THR_VA | 1A2EC26 | uQ0.32 | 0.0063922
72 RESERVED_C72 73 RESERVED_C73 74 AFE_CTRL 75 RESERVED_C75 | [12 empty] | 34 | 1 | 36 | SAG_THR_VB | 1A2EC26 | uQ0.32 | 0.0063922
0 0 E000000 0 | [12 empty] | 34 | 2 | 37 | SAG_THR_VC | 1A2EC26 | uQ0.32 | 0.0063922
76 RESERVED_C76 77 RESERVED_C77 78 PWR_OFFS_P_A 79 PWR_OFFS_P_B | [12 empty] | 34 | 3 | 38 | INT_THR_VA | 68BB0A | uQ0.32 | 0.0015981
0 0 0 0 | [12 empty] | 34 | 4 | 39 | INT_THR_VB | 68BB0A | uQ0.32 | 0.0015981
80 PWR_OFFS_P_C 81 PWR_OFFS_Q_A 82 PWR_OFFS_Q_B 83 PWR_OFFS_Q_C | [12 empty] | 36 | 1 | 40 | INT_THR_VC | 68BB0A | uQ0.32 | 0.0015981
0 0 0 0 | [12 empty] | 36 | 2 | 41 | RESERVED_C41 | 0 | uQ0.32 | 0
>dsr | [12 empty] | 36 | 3 | 42 | RESERVED_C42 | 0 | uQ0.32 | 0
00 VERSION 01 STATUS 02 STATE_FLAG 03 CAPTURE_STATUS | [12 empty] | 36 | 4 | 43 | RESERVED_C43 | 0 | uQ0.32 | 0
30400 8 27 0 | [12 empty] | 38 | 1 | 44 | K_IA | 9A523 | uQ22.10 | 617.2842
04 INTERVAL_NUM 05 N 06 PH_OFFSET 07 FREQ | [12 empty] | 38 | 2 | 45 | K_VA | 19CC00 | uQ22.10 | 1651
1D4000F FA0 FF0A3D71 31FFB | [12 empty] | 38 | 3 | 46 | K_IB | 9A523 | uQ22.10 | 617.2842
08 FREQ_VA 09 FREQ_VB 10 FREQ_VC 11 RESERVED_S11 | [12 empty] | 38 | 4 | 47 | K_VB | 19CC00 | uQ22.10 | 1651
31FFB 31FFB 31FFB 0 | [12 empty] | 40 | 1 | 48 | K_IC | 9A523 | uQ22.10 | 617.2842
12 TEMPERATURE 13 V_A_MAX 14 V_B_MAX 15 V_C_MAX | [12 empty] | 40 | 2 | 49 | K_VC | 19CC00 | uQ22.10 | 1651
0 6093534 608EDAA 6090CD0 | [12 empty] | 40 | 3 | 50 | K_IN | 0 | uQ22.10 | 0
16 RESERVED_S16 17 I_A_MAX 18 I_B_MAX 19 I_C_MAX | [12 empty] | 40 | 4 | 51 | K_VD | 0 | uQ22.10 | 0
0 233A272 23434FF 2371896 | [12 empty] | 42 | 1 | 52 | CAL_M_IA | 20000000 | sQ2.29 | 1
20 I_Nm_MAX 21 I_Ni_MAX 22 FEATURES 23 N_CYCLE | [12 empty] | 42 | 2 | 53 | CAL_M_VA | 20000000 | sQ2.29 | 1
0 58956 807F6E00 50007 | [12 empty] | 42 | 3 | 54 | CAL_M_IB | 20000000 | sQ2.29 | 1
24 RESERVED_S24 25 RESERVED_S25 26 PULSE0_COUNTER 27 PULSE1_COUNTER | [12 empty] | 42 | 4 | 55 | CAL_M_VB | 20000000 | sQ2.29 | 1
0 0 F39 1A65 | [12 empty] | 44 | 1 | 56 | CAL_M_IC | 20000000 | sQ2.29 | 1
28 PULSE2_COUNTER 29 RESERVED_S29 30 ZC_N_VA 31 ZC_N_VB | [12 empty] | 44 | 2 | 57 | CAL_M_VC | 20000000 | sQ2.29 | 1
0 0 54F589 51A02B | [12 empty] | 44 | 3 | 58 | CAL_M_IN | 20000000 | sQ2.29 | 1
32 ZC_N_VC 33 RESERVED_S33 34 ATCAL_41_44 35 ATCAL_45_48 | [12 empty] | 44 | 4 | 59 | CAL_M_VD | 20000000 | sQ2.29 | 1
584AE9 0 0 0 | [12 empty] | 46 | 1 | 60 | CAL_PH_IA | 0 | sQ0.31 | 0
>dar | [12 empty] | 46 | 2 | 61 | CAL_PH_VA | 0 | sQ0.31 | 0
00 V_A 01 V_B 02 V_C 03 V_D | [12 empty] | 46 | 3 | 62 | CAL_PH_IB | 0 | sQ0.31 | 0
45B051F9B426 45AC8C579F3D 45B4A1DD4115 0 | [12 empty] | 46 | 4 | 63 | CAL_PH_VB | 0 | sQ0.31 | 0
04 V_A_F 05 V_B_F 06 V_C_F 07 V_D_F | [12 empty] | 48 | 1 | 64 | CAL_PH_IC | 0 | sQ0.31 | 0
45B0507D2D77 45ABB28F59C6 45B4AC4E3786 0 | [12 empty] | 48 | 2 | 65 | CAL_PH_VC | 0 | sQ0.31 | 0
08 V_AB 09 V_BC 10 V_CA 11 V_AB_F | [12 empty] | 48 | 3 | 66 | CAL_PH_IN | 0 | sQ0.31 | 0
D10A26086387 D117E9ADE66F D1125A16C4E0 D1088CA58502 | [12 empty] | 48 | 4 | 67 | RESERVED_C67 | 0 | uQ0.32 | 0
12 V_BC_F 13 V_CA_F 14 RESERVED_A14 15 RESERVED_A15 | [12 empty] | 50 | 1 | 68 | CPTR_CTRL | 0 | uQ32.0 | 0
D11647A57B48 D11339546E7E 0 0 | [12 empty] | 50 | 2 | 69 | CPTR_BUFF_SIZE | FA0 | uQ32.0 | 4000
16 RESERVED_A16 17 I_A 18 I_B 19 I_C | [12 empty] | 50 | 3 | 70 | CPTR_ADDR | 20000000 | uQ32.0 | 536870912
0 955CCF3945A 9522B8C5887 94F209D5691 | [12 empty] | 50 | 4 | 71 | RESERVED_C71 | 0 | uQ32.0 | 0
20 I_Ni 21 I_Nm 22 I_A_F 23 I_B_F | [12 empty] | 52 | 1 | 72 | RESERVED_C72 | 0 | uQ32.0 | 0
19873E3 0 955C97B4085 9522BE56234 | [12 empty] | 52 | 2 | 73 | RESERVED_C73 | 0 | uQ32.0 | 0
24 I_C_F 25 I_Nmi 26 I_Nm_F 27 RESERVED_A27 | [12 empty] | 52 | 3 | 74 | AFE_CTRL | E000000 | uQ32.0 | 234881024
94F08974F04 19873CD 0 0 | [12 empty] | 52 | 4 | 75 | RESERVED_C75 | 0 | uQ32.0 | 0
28 RESERVED_A28 29 RESERVED_A29 30 RESERVED_A30 31 RESERVED_A31 | [12 empty] | 54 | 1 | 76 | RESERVED_C76 | 0 | uQ32.0 | 0
0 0 0 0 | [12 empty] | 54 | 2 | 77 | RESERVED_C77 | 0 | uQ32.0 | 0
32 P_A 33 P_B 34 P_C 35 P_A_F | [12 empty] | 54 | 3 | 78 | PWR_OFFS_P_A | 0
CC50E80B33E CC2402D57B4 CC1CE955663 CC52E213F6B | [12 empty] | 54 | 4 | 79 | PWR_OFFS_P_B | 0
36 P_B_F 37 P_C_F 38 P_N 39 P_N_F | [12 empty] | 56 | 1 | 80 | PWR_OFFS_P_C | 0
CC210FC5D1C CC1A7B8527A 0 0 | [12 empty] | 56 | 2 | 81 | PWR_OFFS_Q_A | 0
40 RESERVED_A40 41 Q_A 42 Q_B 43 Q_C | [12 empty] | 56 | 3 | 82 | PWR_OFFS_Q_B | 0
0 16143F5C38B6 160F38CA4BFF 160CA21C0654 | [12 empty] | 56 | 4 | 83 | PWR_OFFS_Q_C | 0
44 Q_A_F 45 Q_B_F 46 Q_C_F 47 Q_N
1614065D2D67 160F5AF6DBB8 160C8875505D 0 | [12 empty] | STATUS REGISTERS
48 Q_N_F 49 RESERVED_A49 50 ACC_T0 51 ACC_T1 | [12 empty] | 59 | 1 | 00 | VERSION | 30400 | uQ32.0 | 197632
0 0 130F49B1E3D 210BB84F35F | [12 empty] | 59 | 2 | 01 | STATUS | 8 | uQ32.0 | 8
52 ACC_T2 53 RESERVED_A53 54 RESERVED_A54 | [12 empty] | 59 | 3 | 02 | STATE_FLAG | 27 | uQ32.0 | 39
0 0 0 | [12 empty] | 59 | 4 | 03 | CAPTURE_STATUS | 0 | uQ32.0 | 0
>dpcar | [12 empty] | 61 | 1 | 04 | INTERVAL_NUM | 1D4000F | uQ32.0 | 30670863
00 CYCLE_V_A 01 CYCLE_V_B 02 CYCLE_V_C 03 CYCLE_V_A_F | [12 empty] | 61 | 2 | 05 | N | FA0 | uQ16.0 | 4000
164BFA1E2CD 164D01DDDB2 164E36BBB61 164BB28D2B0 | [12 empty] | 61 | 3 | 06 | PH_OFFSET | FF0A3D71 | sQ0.31 | -0.0075
04 CYCLE_V_B_F 05 CYCLE_V_C_F 06 CYCLE_V_AB 07 CYCLE_V_BC | [12 empty] | 61 | 4 | 07 | FREQ | 31FFB | uQ20.12 | 49.9988
164CE022081 164DF394B25 42E488AE290 42E95DB261C | [12 empty] | 63 | 1 | 08 | FREQ_VA | 31FFB | uQ20.12 | 49.9988
08 CYCLE_V_CA 09 CYCLE_V_AB_F 10 CYCLE_V_BC_F 11 CYCLE_V_CA_F | [12 empty] | 63 | 2 | 09 | FREQ_VB | 31FFB | uQ20.12 | 49.9988
42E7AA4B434 42E3B33443E 42E8DE15527 42E7017EA40 | [12 empty] | 63 | 3 | 10 | FREQ_VC | 31FFB | uQ20.12 | 49.9988
12 CYCLE_I_A 13 CYCLE_I_B 14 CYCLE_I_C 15 CYCLE_I_Nm | [12 empty] | 63 | 4 | 11 | RESERVED_S11 | 0 | uQ32.0 | 0
2FCB8D512F 2FB9A9996D 2FAAB8E613 0 | [12 empty] | 65 | 1 | 12 | TEMPERATURE | 0 | sQ23.8 | 0
16 CYCLE_I_A_F 17 CYCLE_I_B_F 18 CYCLE_I_C_F 19 CYCLE_I_Nm_F | [12 empty] | 65 | 2 | 13 | V_A_MAX | 6093534 | sQ2.29 | 0.18862 | 311.4182
2FCB30DC22 2FB997C042 2FA9970352 0 | [12 empty] | 65 | 3 | 14 | V_B_MAX | 608EDAA | sQ2.29 | 0.18859 | 311.3619
20 CYCLE_P_A 21 CYCLE_P_B 22 CYCLE_P_C 23 CYCLE_P_A_F | [12 empty] | 65 | 4 | 15 | V_C_MAX | 6090CD0 | sQ2.29 | 0.1886 | 311.3864
41640442B4 41547F30D5 415274C285 4164C4F0AF | [12 empty] | 67 | 1 | 16 | RESERVED_S16 | 0 | sQ2.29 | 0
24 CYCLE_P_B_F 25 CYCLE_P_C_F 26 CYCLE_P_N 27 CYCLE_P_N_F | [12 empty] | 67 | 2 | 17 | I_A_MAX | 233A272 | sQ2.29 | 0.068803 | 42.471
4153BF0180 4151FD583D 0 0 | [12 empty] | 67 | 3 | 18 | I_B_MAX | 23434FF | sQ2.29 | 0.068873 | 42.5142
28 CYCLE_Q_A 29 CYCLE_Q_B 30 CYCLE_Q_C 31 CYCLE_Q_A_F | [12 empty] | 67 | 4 | 19 | I_C_MAX | 2371896 | sQ2.29 | 0.069226 | 42.7319
7106F9F3EE 70F6CA88D6 70E41237D6 710559FC5B | [12 empty] | 69 | 1 | 20 | I_Nm_MAX | 0 | sQ2.29 | 0 | 0
32 CYCLE_Q_B_F 33 CYCLE_Q_C_F 34 CYCLE_Q_N 35 CYCLE_Q_N_F | [12 empty] | 69 | 2 | 21 | I_Ni_MAX | 58956 | sQ16.15 | 11.0729 | 11.0729
70F73C3A95 70E236E079 0 | [12 empty] | 69 | 3 | 22 | FEATURES | 807F6E00 | uQ32.0 | 2155834880
[13 empty] | 69 | 4 | 23 | N_CYCLE | 50007 | uQ20.12 | 80.0017
[13 empty] | 71 | 1 | 24 | RESERVED_S24 | 0 | uQ32.0 | 0
[13 empty] | 71 | 2 | 25 | RESERVED_S25 | 0 | sQ2.29 | 0
[13 empty] | 71 | 3 | 26 | PULSE0_COUNTER | F39 | uQ32.0 | 3897
[13 empty] | 71 | 4 | 27 | PULSE1_COUNTER | 1A65 | uQ32.0 | 6757
[13 empty] | 73 | 1 | 28 | PULSE2_COUNTER | 0 | uQ32.0 | 0
[13 empty] | 73 | 2 | 29 | RESERVED_S29 | 0 | uQ32.0 | 0
[13 empty] | 73 | 3 | 30 | ZC_N_VA | 54F589 | uQ20.12 | 1359.3459
[13 empty] | 73 | 4 | 31 | ZC_N_VB | 51A02B | uQ20.12 | 1306.0105 | -53.3354
[13 empty] | 75 | 1 | 32 | ZC_N_VC | 584AE9 | uQ20.12 | 1412.6819 | 53.3359
[13 empty] | 75 | 2 | 33 | RESERVED_S33 | 0 | uQ32.0 | 0
[13 empty] | 75 | 3 | 34 | ATCAL_41_44 | 0 | uQ32.0 | 0
[13 empty] | 75 | 4 | 35 | ATCAL_45_48 | 0 | uQ32.0 | 0
ACCUMULATOR REGISTERS
[13 empty] | 78 | 1 | 00 | V_A | 45B051F9B426 | uQ24.40 | 69.6888 | 217.9207
[13 empty] | 78 | 2 | 01 | V_B | 45AC8C579F3D | uQ24.40 | 69.674 | 217.8976
[13 empty] | 78 | 3 | 02 | V_C | 45B4A1DD4115 | uQ24.40 | 69.7056 | 217.947
[13 empty] | 78 | 4 | 03 | V_D | 0 | uQ24.40 | 0 | 0
[13 empty] | 80 | 1 | 04 | V_A_F | 45B0507D2D77 | uQ24.40 | 69.6887 | 217.9206
[13 empty] | 80 | 2 | 05 | V_B_F | 45ABB28F59C6 | uQ24.40 | 69.6707 | 217.8924
[13 empty] | 80 | 3 | 06 | V_C_F | 45B4AC4E3786 | uQ24.40 | 69.7058 | 217.9472
[13 empty] | 80 | 4 | 07 | V_D_F | 0 | uQ24.40 | 0 | 0
[13 empty] | 82 | 1 | 08 | V_AB | D10A26086387 | uQ24.40 | 209.0396 | 377.4256 | 119.9979 | Law of cosines
[13 empty] | 82 | 2 | 09 | V_BC | D117E9ADE66F | uQ24.40 | 209.0934 | 377.4742
[13 empty] | 82 | 3 | 10 | V_CA | D1125A16C4E0 | uQ24.40 | 209.0717 | 377.4546 | 119.9906 | Law of cosines
[13 empty] | 82 | 4 | 11 | V_AB_F | D1088CA58502 | uQ24.40 | 209.0334 | 377.42 | 119.9973 | Law of cosines
[13 empty] | 84 | 1 | 12 | V_BC_F | D11647A57B48 | uQ24.40 | 209.087 | 377.4684
[13 empty] | 84 | 2 | 13 | V_CA_F | D11339546E7E | uQ24.40 | 209.0751 | 377.4576 | 119.9921 | Law of cosines
[13 empty] | 84 | 3 | 14 | RESERVED_A14 | 0 | uQ24.40 | 0
[13 empty] | 84 | 4 | 15 | RESERVED_A15 | 0 | SQ23.40 | 0
[13 empty] | 86 | 1 | 16 | RESERVED_A16 | 0 | uQ24.40 | 0
[13 empty] | 86 | 2 | 17 | I_A | 955CCF3945A | uQ24.40 | 9.3352 | 29.8206
[13 empty] | 86 | 3 | 18 | I_B | 9522B8C5887 | uQ24.40 | 9.321 | 29.7979
[13 empty] | 86 | 4 | 19 | I_C | 94F209D5691 | SQ23.40 | 9.3091 | 29.7789
[13 empty] | 88 | 1 | 20 | I_Ni | 19873E3 | uQ44.20 | 25.5283 | 0.079888
[13 empty] | 88 | 2 | 21 | I_Nm | 0 | uQ24.40 | 0 | 0
[13 empty] | 88 | 3 | 22 | I_A_F | 955C97B4085 | uQ24.40 | 9.3351 | 29.8205
[13 empty] | 88 | 4 | 23 | I_B_F | 9522BE56234 | uQ24.40 | 9.321 | 29.7979
[13 empty] | 90 | 1 | 24 | I_C_F | 94F08974F04 | uQ24.40 | 9.3087 | 29.7783
[13 empty] | 90 | 2 | 25 | I_Nmi | 19873CD | uQ44.20 | 25.5283 | 0.079888
[13 empty] | 90 | 3 | 26 | I_Nm_F | 0 | uQ24.40 | 0 | 0
[13 empty] | 90 | 4 | 27 | RESERVED_A27 | 0 | SQ23.40 | 0
[13 empty] | 92 | 1 | 28 | RESERVED_A28 | 0 | SQ23.40 | 0
[13 empty] | 92 | 2 | 29 | RESERVED_A29 | 0 | sQ33.30 | 0
[13 empty] | 92 | 3 | 30 | RESERVED_A30 | 0 | sQ33.30 | 0
[13 empty] | 92 | 4 | 31 | RESERVED_A31 | 0 | sQ33.30 | 0
[13 empty] | 94 | 1 | 32 | P_A | CC50E80B33E | SQ23.40 | 12.7698 | 3253.5292
[13 empty] | 94 | 2 | 33 | P_B | CC2402D57B4 | SQ23.40 | 12.7588 | 3250.7366
[13 empty] | 94 | 3 | 34 | P_C | CC1CE955663 | SQ23.40 | 12.7571 | 3250.295
[13 empty] | 94 | 4 | 35 | P_A_F | CC52E213F6B | SQ23.40 | 12.7702 | 3253.6522
[13 empty] | 96 | 1 | 36 | P_B_F | CC210FC5D1C | SQ23.40 | 12.7581 | 3250.5531
[13 empty] | 96 | 2 | 37 | P_C_F | CC1A7B8527A | SQ23.40 | 12.7565 | 3250.1439
[13 empty] | 96 | 3 | 38 | P_N | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
[13 empty] | 96 | 4 | 39 | P_N_F | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
[13 empty] | 98 | 1 | 40 | RESERVED_A40 | 0 | SQ23.40 | 0
[13 empty] | 98 | 2 | 41 | Q_A | 16143F5C38B6 | SQ23.40 | 22.0791 | 5625.4003
[13 empty] | 98 | 3 | 42 | Q_B | 160F38CA4BFF | SQ23.40 | 22.0595 | 5620.3985
[13 empty] | 98 | 4 | 43 | Q_C | 160CA21C0654 | SQ23.40 | 22.0493 | 5617.8222
[13 empty] | 100 | 1 | 44 | Q_A_F | 1614065D2D67 | SQ23.40 | 22.0782 | 5625.1787
[13 empty] | 100 | 2 | 45 | Q_B_F | 160F5AF6DBB8 | SQ23.40 | 22.06 | 5620.5314
[13 empty] | 100 | 3 | 46 | Q_C_F | 160C8875505D | SQ23.40 | 22.049 | 5617.7225
[13 empty] | 100 | 4 | 47 | Q_N | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
[13 empty] | 102 | 1 | 48 | Q_N_F | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
[13 empty] | 102 | 2 | 49 | RESERVED_A49 | 0 | SQ23.40 | 0
0 | [12 empty] | 102 | 3 | 50 | ACC_T0 | 130F49B1E3D | SQ33.30 | 1219.822
[13 empty] | 102 | 4 | 51 | ACC_T1 | 210BB84F35F | SQ33.30 | 2114.93
[13 empty] | 104 | 1 | 52 | ACC_T2 | 0 | SQ33.30 | 0
[13 empty] | 104 | 2 | 53 | RESERVED_A53 | 0 | SQ23.40 | 0
[13 empty] | 104 | 3 | 54 | RESERVED_A54 | 0 | SQ23.40 | 0
PER-CYCLEACCUMULATOR REGISTERS
[13 empty] | 107 | 1 | 00 | CYCLE_V_A | 164BFA1E2CD | UQ24.40 | 1.3935 | 217.9007
[13 empty] | 107 | 2 | 01 | CYCLE_V_B | 164D01DDDB2 | UQ24.40 | 1.3938 | 217.9203
[13 empty] | 107 | 3 | 02 | CYCLE_V_C | 164E36BBB61 | UQ24.40 | 1.3941 | 217.9434
[13 empty] | 107 | 4 | 03 | CYCLE_V_A_F | 164BB28D2B0 | UQ24.40 | 1.3935 | 217.8953
[13 empty] | 109 | 1 | 04 | CYCLE_V_B_F | 164CE022081 | UQ24.40 | 1.3938 | 217.9178
[13 empty] | 109 | 2 | 05 | CYCLE_V_C_F | 164DF394B25 | UQ24.40 | 1.394 | 217.9384
[13 empty] | 109 | 3 | 06 | CYCLE_V_AB | 42E488AE290 | UQ24.40 | 4.1808 | 377.4217
[13 empty] | 109 | 4 | 07 | CYCLE_V_BC | 42E95DB261C | UQ24.40 | 4.182 | 377.4749
[13 empty] | 111 | 1 | 08 | CYCLE_V_CA | 42E7AA4B434 | UQ24.40 | 4.1816 | 377.4562
[13 empty] | 111 | 2 | 09 | CYCLE_V_AB_F | 42E3B33443E | UQ24.40 | 4.1806 | 377.4125
[13 empty] | 111 | 3 | 10 | CYCLE_V_BC_F | 42E8DE15527 | UQ24.40 | 4.1819 | 377.4694
[13 empty] | 111 | 4 | 11 | CYCLE_V_CA_F | 42E7017EA40 | UQ24.40 | 4.1814 | 377.4489
[13 empty] | 113 | 1 | 12 | CYCLE_I_A | 2FCB8D512F | UQ24.40 | 0.1867 | 29.82
[13 empty] | 113 | 2 | 13 | CYCLE_I_B | 2FB9A9996D | UQ24.40 | 0.18643 | 29.7982
[13 empty] | 113 | 3 | 14 | CYCLE_I_C | 2FAAB8E613 | UQ24.40 | 0.1862 | 29.7799
[13 empty] | 113 | 4 | 15 | CYCLE_I_Nm | 0 | UQ24.40 | 0 | 0
[13 empty] | 115 | 1 | 16 | CYCLE_I_A_F | 2FCB30DC22 | UQ24.40 | 0.18669 | 29.8195
[13 empty] | 115 | 2 | 17 | CYCLE_I_B_F | 2FB997C042 | UQ24.40 | 0.18643 | 29.7981
[13 empty] | 115 | 3 | 18 | CYCLE_I_C_F | 2FA9970352 | UQ24.40 | 0.18618 | 29.7786
[13 empty] | 115 | 4 | 19 | CYCLE_I_Nm_F | 0 | UQ24.40 | 0 | 0
[13 empty] | 117 | 1 | 20 | CYCLE_P_A | 41640442B4 | SQ23.40 | 0.25543 | 3253.9353
dsr | [12 empty] | 117 | 2 | 21 | CYCLE_P_B | 41547F30D5 | SQ23.40 | 0.2552 | 3250.9185
00 VERSION 01 STATUS 02 STATE_FLAG 03 CAPTURE_STATUS | [12 empty] | 117 | 3 | 22 | CYCLE_P_C | 415274C285 | SQ23.40 | 0.25516 | 3250.5218
20000 8 16000021 0 | [12 empty] | 117 | 4 | 23 | CYCLE_P_A_F | 4164C4F0AF | SQ23.40 | 0.25544 | 3254.0816
04 INTERVAL_NUM 05 N 06 PH_OFFSET 07 FREQ | [12 empty] | 119 | 1 | 24 | CYCLE_P_B_F | 4153BF0180 | SQ23.40 | 0.25518 | 3250.7726
F760000 FA0 FEC16442 3C000 | [12 empty] | 119 | 2 | 25 | CYCLE_P_C_F | 4151FD583D | SQ23.40 | 0.25516 | 3250.4312
08 FREQ_VA 09 FREQ_VB 10 FREQ_VC 11 RESERVED | [12 empty] | 119 | 3 | 26 | CYCLE_P_N | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
3C000 3BC9A 3BA2A 43 | [12 empty] | 119 | 4 | 27 | CYCLE_P_N_F | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
12 TEMPERATURE 13 I_A_MAX 14 I_B_MAX 15 I_C_MAX | [12 empty] | 121 | 1 | 28 | CYCLE_Q_A | 7106F9F3EE | SQ23.40 | 0.44151 | 5624.3994
3E1 546896B 13B9C 13D0C | [12 empty] | 121 | 2 | 29 | CYCLE_Q_B | 70F6CA88D6 | SQ23.40 | 0.44127 | 5621.2534
16 I_Ni_MAX 17 I_Nm_MAX 18 V_A_MAX 19 V_B_MAX | [12 empty] | 121 | 3 | 30 | CYCLE_Q_C | 70E41237D6 | SQ23.40 | 0.44098 | 5617.6146
3596874 1421A0 121B1C05 15217 | [12 empty] | 121 | 4 | 31 | CYCLE_Q_A_F | 710559FC5B | SQ23.40 | 0.44149 | 5624.0836
20 V_C_MAX 21 RESERVED 22 RESERVED 23 RESERVED | [12 empty] | 123 | 1 | 32 | CYCLE_Q_B_F | 70F73C3A95 | SQ23.40 | 0.44127 | 5621.3397
130DE 12259EF4 12291ABC 1229E218 | [12 empty] | 123 | 2 | 33 | CYCLE_Q_C_F | 70E236E079 | SQ23.40 | 0.44095 | 5617.2536
24 RESERVED 25 RESERVED 26 RESERVED 27 RESERVED | [12 empty] | 123 | 3 | 34 | CYCLE_Q_N | 0 | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
122EE428 121B5760 121B5760 121AB1FC | [12 empty] | 123 | 4 | 35 | CYCLE_Q_N_F |  | SQ23.40 | 0 | 0 | V_A assumed as the reference phase
28 RESERVED 29 RESERVED 30 ZC_N_VA 31 ZC_N_VB
121AFD54 121A6423 F9F115 F6B4CD
32 ZC_N_VC 33 AT_CAL_41_44 34 AT_CAL_45_48
F53D8A 8601CA03 FA01C003

## TOOLS
TOOLS
CALCULATORS
STEP 1: READ THE BOARD CONFIGURATION AND THE SYSTEM FREQUENCY
Send a "DCR" command, read the results and copy it in the proper section of step 4 in the "Step by step guide" sheet
Update the system frequency and update it in the proper section of step 1 in the "Step by step guide" sheet
STEP 2: SET THE DESIRED VALUES
In the yellow cells, type the values to be configured
PARAMETER | NEW VALUE | DEFAULT VALUE | NEEDED COMMANDS | NOTES | PARAMETER 1 | VALUE | PARAMETER 2 | VALUE | PARAMETER 3 | VALUE
Pulse0_Kt | 3200 | 3200 | DCW[9](500000) | Check if the pulse width allows to reach the expected max pulse rate (max power applied)
Pulse1_Kt | 3200 | 3200 | DCW[11](500000)
Pulse2_Kt | 3200 | 3200 | DCW[13](500000)
Creep threshold (P) | 11 | 0 | DCW[15](00010052) | Enable/Disable with Feature Control Register
Creep threshold (Q) | 11 | 0 | DCW[19](00010052)
Creep threshold (I) | 10 | 0 | DCW[23](0000425B)
Creep threshold (S) | 11 | 0 | DCW[27](00010052)
Power offset (PA) | 375 | 0 | DCW[78](E7E2AD85) | Enable/Disable with power offset control register
Power offset (PB) | 5 | 0 | DCW[79](FFADB065)
Power offset (PC) | 1 | 0 | DCW[80](FFEF89AE)
Power offset (QA) | -5 | 0 | DCW[81](00524F9B)
Power offset (QB) | 5 | 0 | DCW[82](FFADB065)
Power offset (QC) | 1 | 0 | DCW[83](FFEF89AE)
Global P offset | 0.05 | 0 | DCW[29](FFFFFED6)
Global Q offset | -0.05 | 0 | DCW[30](0000012A)
Global S offset | -0.05 | 0 | DCW[31](0000012A)
K_VA | NA | 19CC00 | DCW[45](0019CC00) |  | Resistor divider ratio | 1651
K_VB | NA | 19CC00 | DCW47](0019CC00) |  | Resistor divider ratio | 1651
K_VC | NA | 19CC00 | DCW[49](0019CC00) |  | Resistor divider ratio | 1651
K_IA | NA | 9A523 | DCW[44](0004D291) |  | CT Ratio | 1000 | Burden resistor | 1.62 | PGA | 1
|  |  | DCW[44](00138800) |  | Shunt resistor | 100 | PGA | 8
|  |  | DCW[44](0007D000) |  | Rogowski coil scale factor at 60Hz | 500 | PGA | 4
K_IB | NA | 9A523 | DCW[46](0009A523) |  | CT Ratio | 2000 | Burden resistor | 1.62 | PGA | 1
|  |  | DCW[46](00138800) |  | Shunt resistor | 100 | PGA | 8
|  |  | DCW[46](0007D000) |  | Rogowski coil scale factor at 60Hz | 500 | PGA | 4
K_IC | NA | 9A523 | DCW[48](0009A523) |  | CT Ratio | 2000 | Burden resistor | 1.62 | PGA | 1
|  |  | DCW[48](00138800) |  | Shunt resistor | 100 | PGA | 8
|  |  | DCW[48](0007D000) |  | Rogowski coil scale factor at 60Hz | 500 | PGA | 4
K_IN | NA | 9A523 | DCW[50](0009A523) |  | CT Ratio | 2000 | Burden resistor | 1.62 | PGA | 1
|  |  | DCW[50](00138800) |  | Shunt resistor | 100 | PGA | 8
|  |  | DCW[50](0007D000) |  | Rogowski coil scale factor at 60Hz | 500 | PGA | 4
PULSE RATE CALCULATOR (USEFUL TO CONFIGURE THE PULSE WIDTH)
Enter the values in the yellow cells. The blue cell shows the result.
METER CONSTANT | APPLIED POWER | PULSE RATE
3200 | 11.25 | 10

## Version control
Version | Date | Changes
6 | 2025-12-12 | Register set changed. It is valid for It is valid for PIC32CXMTx metrology library3.04.00 or newer
5.2 | 2025-03-07 | Fixed a bug in Step 3 of "Step by Step Guide" (the ordinal of the registers was not correct)
5.1 | 2024-11-13 | Register set changed It is valid for PIC32CXMTx metrology library3.03.00 or newer Valid PIC32CXMTx. SAM4CMx is not supported.
5 | 2023-09-07 | "Tools" worksheet added
4 | 2022-10-19 | This version supports the new features of the neutral current measurements. It requires SAM4CM metrology library 2.0.0B or newer; or PIC32CXMTx 3.0.0A or newer Valid for SAM4CM and PIC32CX
3.1 | 2022-07-29 | Fixed issue calculating CAL_PH_Ix when "Meter Test Freq" (Q5) is different from 60Hz. Calculations of P_x and Q_X quantities fixed (a constant of "4000" was used instead of status register "N")
3 | 2022-02-22 | Initial edition